|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Product Name** | [Minuet](#Minuet)http://www.arjohuntleigh.com.au/PageFiles/1448/minuet.jpg | [Arminia II](#Arminia) https://encrypted-tbn3.gstatic.com/images?q=tbn:ANd9GcTu9CI411GtPJnghBF386XBvFlQky-60ZG8V-KnLZ041Ih1yXO4Rw | MelanieVerstellung-Melanie | [Etude Plus](#Etude)http://www.gmsrehab.com.au/images/ETPLUSTHUMB.jpg |
| Manufacturer | Arjo Huntleigh | Burmeir | Betten Malsch | Invacare |
| Availability through DES | **refurbished only** | **refurbished only** | **purchase\* or refurbished** | **purchase\* or refurbished** |
| DES Code | AB01I A | AB01O A | AB01P A | AB01T A |
| SWL capacityNB: allow approx. 40kg for mattress & bedding | 180kg | 175kg | 175kg | 200kg |
| Features | head raise, knee bend,manual leg raiseTrendelenburgreverse Trendelenburg | head raise, knee bend,leg raise manual, Trendelenburg (some have reverse Trendelenburg) | head raise, knee bend,leg raise manual, Trendelenburgreverse Trendelenburg | head raise, knee bend,leg raise manual, Trendelenburgreverse Trendelenburg |
| Total bed footprint (W x L) | 920 x 2200 | 1030 x 2200 | 1050 x 2130 | 1020 x 2200 |
| Height range | 220 - 630 | 400 - 800 | 400 - 800 | (standard) 410 – 810 OR 340 -740  |
| Clearance under bed | clear | clear | clear | clear |
| Backrest platform length | 760 | 750 | 700 | 850 |
| Backrest angle | 0° - 73° | 0° - 70° |   | 0° - 70° |
| Internal platform dimensions (W x L)  | 920 x 2040 | 890 x 2050 | 890 x 2000 | 890 x 2080 |
| Lockable handle controller? | no | no | yes | optional |
| Mattress retainers? | Side (head end)Foot end | Sides | no | yes, can rotate away |
| Width between retainers | 870 (width)2000 (length) | 850 | n/a | 850 |
| Bed extension available**?** | no | yes – 150mm | no | yes – 200mm |
| Self-help pole position | 90mm in from headboard70mm from side railing | 75mm in from headboard70mm from side railing | n/a | 80mm in from headboard 60mm from side railing |
| Other notes |  |  |  | Not suitable for <12yo |

|  |
| --- |
| **Minuet** |
| **Bed Rail Options and Specifications** |
| Style | 3/4 length | **C:\temp\ViewingImage.jpg** |
| DES Code | AB09G A |
| Internal gaps within rails  | 120 |
| Distance between left and right bed rails | 980 |
| Rail height from bed platform | 395 |
| Rail cover options | Padded – ABO9I A | C:\temp\ViewingImage.jpg |
| **Information on Mattress Requirements** |  |
| **Configuration** | **Max. Height** | **Max. Width** (no side retainers) | **Max. Width** (side retainers) | **Max. Length** |
| no rails | n/a | 920 | 870 | 2000 |
| 3/4 rails | 175 | 940 | 870 | 2000 |
| **Potential Suitable Mattresses** | **NB: does not indicate entrapment zone compliance** |
| **Product Name** | **DES Code** | **Equip. Program Category** | **Mattress Dimensions** **(W x L x H)** | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | AB14J A | Cat 1 | 870 x 1960 X 150 | ¾ rails |
| Zenith PM265 Mattress Overlay | AB15AA A | Cat 1 | 900 x 2000 x 127 (+100 mm underlay) | no rails |
| Zenith PM268 Mattress Replacement System | AB15AB A | Cat 1 | 900 x 2000 x 200 | no rails |

|  |
| --- |
| **Arminia II** |
| **Bed Rail Options and Specifications** |
| Style | Full length |  |
| DES Code | n/a (ordered with bed) |
| Internal gaps within rails  | 150 |
| Distance between left and right bed rails | 940 |
| Rail height from bed platform | 385 |
| Rail cover options | full length - AB09N A | C:\temp\ViewingImage.jpg |
| **Information on Mattress Requirements** |  |
| **Configuration** | **Max. Height** | **Max. Width**(no side retainers) | **Max. Width** (side retainers) | **Max. Length** |
| no rails | n/a | 900 | 850 | 2010 |
| Full length rails | 165 | 900 | 850 | 2010 |
| **Potential Suitable Mattresses** | **NB: does not indicate entrapment zone compliance** |
| **Product Name** | **DES Code** | **Equip. Program Category** | **Mattress Dimensions** (W x L x H) | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | AB14J A | Cat 1 | 870 x 1960 X 150 | Full length |
| Mattress ROHO 4 Sections Includes Cover | AB15F A | Cat 1 | 860 x 2030 x 80 | Full length |
| Zenith PM265 Mattress Overlay | AB15AA A | Cat 1 | 900 x 2000 x 127 (+75mm underlay) | Full length |
| Zenith PM268 Mattress Replacement System | AB15AB A | Cat 1 | 900 x 2000 x 200 | Full length |

|  |
| --- |
| **Etude Plus** |
| **Bed Rail Options and Specifications** |
|  | **Scala Basic 2** | **Full length** |
| **C:\temp\ViewingImage.jpg** | **C:\Users\larthu01\Desktop\ESIS images\20150820_143524.jpg** |
| DES Code | AB09J A | AB09F AAB09Q A (extended length rails) |
| Internal gaps within rails | 100 | 100 |
| Distance between left and right bed rails | 910 | 955 - 970 |
| Rail height from bed platform | 390 | platform at lowest setting: 440platform at highest setting: 370 |
| Rail cover options | C:\Users\larthu01\Desktop\ESIS images\20150820_150551.jpg Padded – AB09K A | C:\Users\larthu01\Desktop\ESIS images\20150820_143713.jpgMesh - AB09O A; AB09S A (for extended length rails) |
|  | Padded - AB09P A; AB09R A (for extended length rails)C:\Users\larthu01\Desktop\ESIS images\20150820_141651.jpg |
| **Information on Mattress Requirements**  |  |
| Configuration | **Max. Height** | **Max. Width**(no side retainers) | **Max. Width** (with side retainers) | **Max. Length** |
| no rails | n/a | 900 | 850 | 2040 |
| Scala Basic 2 | 170 | 880 | 850 | 2040 |
| Full length rails | lowest setting: 220highest setting: 150 | 900 | 850 | 2040 |
| **Potential Suitable Mattresses**  | **NB: does not indicate entrapment zone compliance** |
| **Product Name** | **DES Code** | **Equip. Program Category** | **Mattress Dimensions** (W x L x H) | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | AB14J A | Cat 1 | 870 x 1960 X 150 | all |
|  Zenith foam mattress - narrow | AB14O A | Cat 1 | 840 x 2020 x 150 | all |
| Mattress Platinum Foam | AB14Q A | Cat 1 | 880 x 2020/2040 x 150 | all |
| Mattress ROHO 4 Sections Includes Cover | AB15F A | Cat 1 | 860 x 2030 x 80 | all |
| Zenith PM265 Mattress Overlay | AB15AA A | Cat 1 | 900 x 2000 x 127 (+75mm underlay) | Full length |
| Zenith PM268 Mattress Replacement System | AB15AB A | Cat 1 | 900 x 2000 x 200 | Full length |
| Mattress Overlay AlphaXcell | AB15T A | Cat 2 | 840 x 2050 x 120 | no rails |